GAS CHROMATOGRAPH  «CRYSTALLUX-4000М»

Dear Sirs! To call our specialist for the Сristallux-4000M chromatograph commissioning, please fill out the form below and send it by e-mail: bashirova@meta-chrom.ru , meta-ola@yandex.ru , m_chrom@mari-el.ru (please be sure to include your address, phone number or e-mail, position and full name of the contact person for communication).

QUESTIONNAIRE

We inform you that the preparatory work in accordance with the Operation Manual MCUB.415338.001OM for the installation and commissioning of the chromatograph, received under contract number _____________ with RPC "META-CHROM", CO. LTD, has been fully completed:

- Provided with tables for placing a chromatograph with dimensions of 1800 x 750 mm (paragraph 9.3).

- The floor and tables can support the weight of the 90 kg devices (paragraph 9.3).

- Provided with at least 6 Eurostandard power sockets with grounding and in accordance with the rules of electrical safety (paragraph 9.2).

- The AC network meets the requirements of 220 V (+ 10 / -15%), 50 Hz, 1.5 kW and is free from slow and sharp changes in the average voltage value and impulse noise (paragraph 9.2).

- The workplace is provided with the necessary gases of the required purity, gas lines are installed (if the installation of gas lines by the installer is not specified in the contract). Provided with all the necessary reducers, deionized water (0.8 Mohm / cm) for the hydrogen generator (paragraph 9.5).

- The workplace in terms of temperature, humidity and air purity, illumination of the workplace meets the sanitary and hygienic requirements (paragraph 2.2).

- The complete set of delivery of the chromatograph is available at the installation site.

- The packing of the chromatograph and accessories is not broken.

- the operator of the chromatograph, who will work on it in the future, must be present during the entire period of commissioning works of the chromatograph

- A personal computer was purchased and installed at the workplace that meets the requirements (if it is not included in the delivery set),

- Provided with standard samples or calibration mixtures for the calibration of the chromatograph (if not included in the delivery set).

Please send a specialist to carry out commissioning works at:_________________________________________________________________________

Name of the organization ____________________________________________ Phone number ___________ Contact persons __________________________________________________________  Preferred arrival date of the specialist _________________________

We acknowledge that incomplete preparation on our part for the launch of the equipment may cause additional costs or a repeated visit, which we undertake to pay according to the invoice (additional time spent by the specialist, additional transport and travel expenses, hotel payment).

Date_________________		                    Signature _____________________